## Slide 1: DS22 noise studies
Nikolai Beev
TE-EPC-HPM
26.04.2018

## Slide 2: HL-LHC Requirements
Document: CERN-ACC-2017-0101
http://cds.cern.ch/record/2298764/files/CERN-ACC-2017-0101.pdf
Page 15, table 6
LF noise, machine temperature cycle
LF noise, toughest specification
Broadband noise.
Mostly for monitoring, but its LF part may be significant for regulation up to a few Hz.
LF noise, temperature cycle
50:50 split with safety margin:
0.2 ppm (2xRMS) -> 0.1 ppmRMS
ADC
DCCT
0.05 ppmRMS
500 nVRMS
RTI, FS = 10 V
0.05 ppmRMS
500 nVRMS
RTO, FS = 10 V

## Slide 3: Noise Measurements
Working in the frequency domain has numerous advantages:
the type of noise is immediately visible (white, 1/f, narrowband peaks, etc.)
Bandwidths of interest can be defined sharply (e.g. 0.001-0.1 Hz)
HL-LHC fill stability and short-term stability requirements are defined through frequency bands

## Slide 4: A bit on FFT and PSD estimation
In order to resolve LF components in the frequency domain, the time-domain record has to be sufficiently long
For the DS22, we use records of 1600 s (≈26 min). The first resolved FFT bin is 1.25 mHz
A single FFT estimate (periodogram) has high variance. We decrease it by averaging multiple windowed periodograms (Bartlett method)
Typically, we average >30 FFTs from an overnight measurement cycle (>12 h)

## Slide 5: An example
random white noise
x window (Hanning)
Modified periodogram
Bartlett, 10x averages
Welch, 20x averages
FFT length is set by the window length
FFT,
Averaging of FFTs

## Slide 6: Unmodified DS22 Noise
modulator
quantization noise…
well understood,
largely unimportant
1/f region
chopper
stage /
crossover
white
noise
well understood
EDMS 1792151 v.1
UNNOWN
until 2018
Poorly known
until 2018
EDMS 1953928 v.1
well understood

## Slide 7: DS22 improvements
HL-LHC Class 0
short-term stability requirement BW
605 nVRMS
220 nVRMS
≈120 nVRMS
--------------------------
in total: 5x improvement
2.3 effective bits
Noise behaviour for f >0.01 Hz
very well understood
(LF, crossover, HF)

## Slide 8: Dominant noise source
By replacing op amps one by one, it was determined that IC10 contributes almost all of the 1/f noise. The origin is most likely its 1/f current noise.

## Slide 9: Other Culprits
When the DS22 box is open, higher LF noise is observed due to free air flow (TC / thermal EMFs?)

## Slide 10: Reference noise
With shorted inputs, the noise originating from Vref is suppressed due to the symmetry of the +5V and -5V ref. lines
Vref noise is also suppressed when the DS22 measures its own Vref
+10V or -10V from a PBC show the same LF noise level

## Slide 11: LTZ1000 LF noise
GOOD NEWS:
Even considering the
worst-case scenario,
at +10 V input we are still
safely within the HL-LHC specs
BAD NEWS:
We can’t measure reliably
below this line.
PBC noise is on the same order
≈140-160 nVRMS (0.001-0.1 Hz)
200 nVRMS
120 nVRMS
≈160 nVRMS uncertain
PBC or DS22 reference?
Device-to-device variation?

## Slide 12: More on LTZ1000 LF noise
We would like to better understand the LF noise of LTZ1000
For all of the new ADCs, we should characterize and keep track of the LTZs
Study LF noise before and after burn-in?
Ultimate metrological characterization: ADC against JJ calibrator. Collaboration with METAS/PTB/etc. foreseen.

## Slide 13: Things to improve
Input buffer, Vref buffers (OPA227, OPA277) - to be replaced with ADA4522-1
All thick-film resistors in the signal / Vref path to be replaced with thin-film
All non-NP0 ceramic capacitors in the signal / Vref path to be replaced with plastic (ECPU)
Small improvement (≈2 dB) in broadband noise is possible with some filtering of the -5V ref.
Temperature stabilization loop can be improved (work in progress)

## Slide 14: Conclusions
DS22 can be improved considerably in terms of LF noise
The HF noise of DS22 is presently dominated by the digital filter / decimation in the FPGA. There is no strong pressure to improve it in the DS22 hardware. By itself it meets the specs.
Ultimately, at very LF (< 0.01 Hz) any ADC will be limited by the reference noise (LTZ1000)
We are at the limit of our measurement capability when it comes to LTZ1000 LF noise