## Sheet
PERFORMANCE TEST
DUT
Manufacturer |  | Datron Instruments |  |  | S/N |  | 27316
Model Number |  | 4950 |  |  | Test date |  | 19-06-28 01:23:53
Meter Info |  | Datron 4950 STD |  |  | Ambient Temp. |  | 23.0
Options |  | 1,1,1 |  |  | CERT_AMB? |  | 23
Firmware |  | 890197-02.08 |  |  | TEMP? |  | 40.73
Date of Baseline Cal |  | "2017-MAR-26" |  |  | TEMP? BASE |  | 4.2
Date of Certified Cal |  | "2019-JUN-18" |  |  | TEMP? CERTIFIED |  | 0.69
CNFTST? |  | 0 |  |  | STANDBY? |  | "01:45"
DUT Confidence level |  | 30 Days Transfer Satbility
Calibrator
Reference standard |  | Options |  | Unc | Calibration Date |  | Due Date
Fluke 5700A EP |  | 03 WB |  | -- | -- |  | --
Fluke 732A#2 |  | 10.000000 VDC |  | -- | -- |  | --
IET SRL-1 |  | 1.0000678 ohm |  | -- | -- |  | --
ESI SR-104 |  | 10.0000125 kohm |  | 0.3ppm | 12/18/2015 |  | 12/18/2016
MFC
MFC Calibrate Date |  | 230319 |  |  | MFC last calibrated |  | 96 days
Calibrate Date WBFLAT |  | 070999 |  |  | MFC since WBFLAT |  | 0 days
Calibrate Date WB Gain |  | 030505 |  |  | MFC since WBGAIN |  | 5168 days
MFC Calibrate Date Zero |  | 150619 |  |  | MFC since DCV Zero |  | 12 days
MFC Confidence level |  | 5700A 99% 90 Days |  |  | CAL CONST 11V Range +zero |  | 398.1926
CAL CONST 6.5V |  | +6.917987472993364E+00 |  |  | CAL CONST 11V Range -zero |  | 398.1921
CAL CONST 13V reference voltage |  | +1.379989402253506E+01 |  |  | CAL CONST, Zero calibration temperature |  | 23.5
CAL CONST 10KOHM true output resistance |  | +9.999939655173794E+03 |  |  | CAL CONST, ALL calibration temp |  | 23.5
CAL SONST 10kOHM standard resistance |  | +9.999122357832126E+03 |  |  | CAL_CONST? CHECK, WB_FLAT_TEMP |  | 23
CAL CONST 1 OHM |  | +9.998277603537725E-01 |  |  | CAL CONST,WBAND_TEMP
ETIME? |  | 967 days 11.82 hours |  |  | CAL TEMP? CHECK
ONTIME? |  | 0 days 4.25 hours |  |  | CAL CONST,DACLIN |  | 0
DCV PERFORMANCE TEST
DCV Test VDC | Expected Value | Measured μV | Source Unc. μV | Lower Limit μV | Upper Limit μV | Deviation | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV μV
0.0mVDC | 0 | -0.17 | 7.5e-07 |  |  | N/A |  |  | 0.015092 | 0.00971
0.0VDC | 0 | 0 | 7.5e-07 |  |  | N/A |  |  | 0.070711 | 0.01484
0.0VDC | 0 | 0 | 7.5e-07 |  |  | N/A |  |  | 0.33333 | 0.0827
0.0VDC | 0 | -10 | 7.5e-07 |  |  | N/A |  |  | 3.3333 | 0.99
0.0VDC | 0 | 100 | 7.5e-07 |  |  | N/A |  |  | 0 | 4.4
DCV Test VDC | DMM Range | Measured V | Source Unc. ppm | Lower Limit V | Upper Limit V | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.01 | 100mV | 0.0099997 | 82 |  |  | -27 |  |  | 1.5093 | 7.3e-07
0.02 | 100mV | 0.02 | 44.5 |  |  | -9.5 |  |  | 0.89366 | 3.9e-07
0.05 | 100mV | 0.05 | 22 |  |  | -2.2 |  |  | 0.45947 | 1.6e-07
0.1 | 100mV | 0.1 | 14.5 | 0.099999 | 0.1 | 0.7 | 3 | 4.7275 | 0.18333 | 7e-08
-0.01 | 100mV | -0.01 | -82 |  |  | -4 |  |  | -2.3688 | 7.6e-07
-0.02 | 100mV | -0.02 | -44.5 |  |  | 1.5 |  |  | -1 | 3.9e-07
-0.05 | 100mV | -0.05 | 22 |  |  | 1.4 |  |  | -0.58973 | 1.5e-07
-0.1 | 100mV | -0.1 | 14.5 | -0.099999 | -0.1 | 2.8 | 3 | 18.9099 | -0.10138 | 8e-08
0.1 | 1V | 0.1 | 14.5 |  |  | 0 |  |  | 1.3944 | 1.6e-07
0.2 | 1V | 0.2 | 10.75 |  |  | 2.5 |  |  | 0.26352 | 8e-08
0.5 | 1V | 0.5 | 8.4 |  |  | 0.8 |  |  | 0.088192 | 5e-08
1 | 1V | 1 | 7.2 | 0.99999 | 1 | 1 | 1.5 | 13.597 | 0.078174 | 4e-08
-0.1 | 1V | -0.1 | 14.5 |  |  | 0 |  |  | -1.2247 | 1.5e-07
-0.2 | 1V | -0.2 | 10.75 |  |  | 2.5 |  |  | -0.26352 | 7e-08
-0.5 | 1V | -0.5 | 8.4 |  |  | 0.8 |  |  | -0.1 | 3e-08
-1 | 1V | -1 | 7.2 | -0.99999 | -1 | 1.4 | 1.5 | 19.0357 | -0.1118 | 2e-08
1 | 10V | 0.99999 | 7.2 |  |  | -5 |  |  | 0.72649 | 9e-08
2 | 10V | 2 | 6.6 |  |  | -2 |  |  | 0.25 | 6e-08
5 | 10V | 5 | 5.8 |  |  | -0.8 |  |  | 0.10541 | 3e-08
10 | 10V | 10 | 5.4 | 9.9999 | 10.0001 | 0 | 1.5 | 0 | 0.083333 | 3e-08
19 | 10V | 19 | 5.4211 | 18.9999 | 19.0001 | 1.1053 | 1.5 | 19.6498 | 0.071801 | 2e-08
-1 | 10V | -1 | 7.2 |  |  | 2 |  |  | -1.7638 | 9e-08
-2 | 10V | -2 | 6.6 |  |  | 1.5 |  |  | -0.25 | 4e-08
-5 | 10V | -5 | 5.8 |  |  | 0.8 |  |  | -0.1 | 2e-08
-10 | 10V | -10 | 5.4 | -9.9999 | -10.0001 | 0.8 | 0.5 | 14.7517 | -0.070711 | 2e-08
-19 | 10V | -19 | 5.4211 | -18.9999 | -19.0001 | 0.94737 | 1.5 | 16.8427 | -0.087719 | 2e-08
10 | 100V | 10 | 5.4 |  |  | -4 |  |  | 1.5899 | 1e-07
20 | 100V | 20 | 5.4 |  |  | -0.5 |  |  | 0.30046 | 7e-08
50 | 100V | 50 | 8 |  |  | -0.8 |  |  | 0 | 4e-08
100 | 100V | 100 | 7 | 99.9993 | 100.0007 | -0.1 | 2 | -1.3736 | 0.072648 | 3e-08
-10 | 100V | -10 | 5.4 |  |  | 3 |  |  | -1.2247 | 1.1e-07
-20 | 100V | -20.0001 | 5.4 |  |  | 3.5 |  |  | -0.25 | 5e-08
-50 | 100V | -50.0001 | 8 |  |  | 2.2 |  |  | -0.088192 | 3e-08
-100 | 100V | -100.0002 | 7 | -99.9993 | -100.0007 | 2.1 | 2 | 28.8457 | -0.13229 | 2e-08
-100 | 1000V | -100.0001 | 7 |  |  | 1 |  |  | -0.72648 | 4e-08
-200 | 1000V | -199.9999 | 6.5 |  |  | -0.5 |  |  | -0.16667 | 3e-08
-500 | 1000V | -500.0004 | 9.2 |  |  | 0.8 |  |  | -0.066667 | 2e-08
-1000 | 1000V | -1000.0017 | 8.6 | -999.9912 | -1000.0088 | 1.7 | 2 | 19.2536 | -0.060092 | 2e-08
100 | 1000V | 99.9994 | 7 |  |  | -6 |  |  | 2 | 8e-08
200 | 1000V | 199.9996 | 6.5 |  |  | -2 |  |  | 0.25 | 5e-08
500 | 1000V | 500.0004 | 9.2 |  |  | 0.8 |  |  | 0 | 3e-08
1000 | 1000V | 1000.0017 | 8.6 | 999.9912 | 1000.0088 | 1.7 | 2 | 19.2536 | 0.1118 | 2e-08
OHM PERFORMANCE TEST
OHM Test 4W | Calibrator ohm | Measured ohm | Source Unc. ppm | Lower Limit ohm | Upper Limit ohm | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 Ω | 0.99983 | 0.99982 | 95 | 0.99973 | 0.99973 | -11.802 | 20 | -12.1567 | 3.3836 | 3.7e-06
1.9 Ω | 1.8998 | 1.8998 | 95 | 1.8996 | 1.9 | -13.8962 | 15 | -14.4485 | 2.2961 | 2.33e-06
10 Ω | 10.0019 | 10.0024 | 28 | 10.0016 | 10.0022 | 49.9906 | 5 | 175.7575 | 2.3957 | 1.64e-06
19 Ω | 19.0022 | 19.003 | 26 | 19.0016 | 19.0028 | 39.7322 | 5 | 150.0665 | 2.4178 | 9.2e-07
100 Ω | 99.9946 | 99.9995 | 17 | 99.9926 | 99.9966 | 48.4026 | 3 | 280.3887 | 2.4579 | 1.58e-06
190 Ω | 189.9811 | 189.9887 | 17 | 189.9773 | 189.9849 | 40.1619 | 3 | 232.6516 | 2.4128 | 8.4e-07
1 kΩ | 999.9256 | 999.9275 | 12 | 999.9106 | 999.9406 | 1.9001 | 3 | 15.3617 | 0 | 3e-08
1.9 kΩ | 1899.6882 | 1899.698 | 12 | 1899.6597 | 1899.7167 | 5.1587 | 3 | 41.706 | 0.10118 | 2e-08
10 kΩ | 9999.94 | 9999.955 | 11 | 9999.8 | 10000.08 | 1.5 | 3 | 13.1559 | 0.052705 | 4e-08
19 kΩ | 18998.837 | 18998.946 | 11 | 18998.571 | 18999.103 | 5.7372 | 3 | 50.3185 | 0.037218 | 3e-08
100 kΩ | 100003.44 | 100002.89 | 13 | 100001.6399 | 100005.2401 | -5.4998 | 5 | -39.4863 | 0.24036 | 3e-08
190 kΩ | 189995.81 | 189995.72 | 13 | 189992.3901 | 189999.2299 | -0.47369 | 5 | -3.4009 | 0.19713 | 2e-08
1 MΩ | 1000033.1 | 999978.1 | 18 | 1000007.0991 | 1000059.1009 | -54.9982 | 8 | -279.211 | 0.91578 | 8e-08
1.9 MΩ | 1899901.5 | 1899824 | 19 | 1899850.2027 | 1899952.7973 | -40.7916 | 8 | -197.8683 | 0.16645 | 6e-08
10MΩ | 9999169 | 9994073 | 37 | 9998679.0407 | 9999658.9593 | -509.6424 | 12 | -1310.2255 | 3.7485 | 2.4e-07
19 MΩ | 18995685 | 18986617 | 47 | 18994564.2546 | 18996805.7454 | -477.3716 | 12 | -984.1144 | 2.5915 | 2.4e-07
100MΩ(2W) | 99992300 | 131910 | 120 | 99962302.31 | 100022297.69 | -998680.7984 | 180 | -461640.3624 | 98456.1836 | 0.001076
OHM Zero 4W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | 4e-06 | 5e-05 |  |  | N/A |  |  | 4.7199e-06 | 3.25e-06
0 Ω | 100 Ω | 0 | 5e-05 |  |  | N/A |  |  | 3.3333e-06 | 4.27e-06
0 Ω | 1 KΩ | -0.0001 | 5e-05 |  |  | N/A |  |  | 7.0711e-05 | 3.16e-05
0 Ω | 10 KΩ | 0 | 5e-05 |  |  | N/A |  |  | 0.00033333 | 0.000201
0 Ω | 100 KΩ | 0 | 5e-05 |  |  | N/A |  |  | 0 | 0.00079
0 Ω | 1 MΩ | -0.1 | 5e-05 |  |  | N/A |  |  | 0 | 0.00902
0 Ω | 10 MΩ | -2 | 5e-05 |  |  | N/A |  |  | 0 | 0.0578
0 Ω | 100 MΩ(2W) | -10 | 5e-05 |  |  | N/A |  |  | 3.3333 | 0.523
OHM Zero 2W | DUT Range | Measured ohm | Source Unc. ohm | Lower Limit ohm | Upper Limit ohm | Measured | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ohm
0 Ω | 10 Ω | -0.00231 | 5e-05 |  |  | N/A |  |  | 3.3208e-06 | 5.58e-06
0 Ω | 100 Ω | -0.00228 | 5e-05 |  |  | N/A |  |  | 8.6603e-06 | 4.44e-06
0 Ω | 1 KΩ | -0.0017 | 5e-05 |  |  | N/A |  |  | 0.00047697 | 2.53e-05
0 Ω | 10 KΩ | 0.001 | 5e-05 |  |  | N/A |  |  | 0.0030732 | 0.000248
0 Ω | 100 KΩ | 0.01 | 5e-05 |  |  | N/A |  |  | 0 | 0.000822
0 Ω | 1 MΩ | 0.1 | 5e-05 |  |  | N/A |  |  | 0 | 0.00858
0 Ω | 10 MΩ | 2 | 5e-05 |  |  | N/A |  |  | 0.5 | 0.05
0 Ω | 100 MΩ | 40 | 5e-05 |  |  | N/A |  |  | 5 | 0.553
ACV PERFORMANCE TEST
ACV Linearity Checks | DUT Range | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm of Input | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1VAC/1kHz | 10 VAC | 1.0007 | 82 | 0.99992 | 1.0001 | 720 | 10 | 871.5915 | 12.0099 | 3.57e-06
2VAC/1kHz | 10 VAC | 2.0007 | 78.5 | 1.9998 | 2.0125 | 335 | 10 | 423.3305 | 2.2041 | 2.16e-06
5VAC/1kHz | 10 VAC | 5.0004 | 89 | 4.9996 | 5.0004 | 72 | 10 | 80.393 | 1.4141 | 1.55e-06
10VAC/1kHz | 10 VAC | 9.9999 | 82 | 9.9992 | 1.0068 | -7 | 10 | -8.4738 | 1.118 | 1.03e-06
12VAC/1kHz | 10 VAC | 11.9997 | 80.833 | 11.999 | 1.0001 | -21.6667 | 10 | -26.6014 | 1.0578 | 1.29e-06
15VAC/1kHz | 10 VAC | 14.9994 | 79.667 | 14.9988 | 1.0064 | -38 | 10 | -47.3272 | 1.4143 | 1.22e-06
19VAC/1kHz | 10 VAC | 18.9991 | 78.684 | 18.9985 | 19.0015 | -47.8947 | 10 | -60.384 | 1.1471 | 1.05e-06
ACV Performance Tset |  | Measured Vac | Source Unc. ppm | Lower Limit Vac | Upper Limit Vac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
0.001 V, 10 Hz |  | 0.0010103 | 5550 | 0.00099445 | 0.0010056 | 10350 | 20 | 186.4853 | 1084.3188 | 0.000233
0.001 V, 20 Hz |  | 0.0010146 | 5220 | 0.00099478 | 0.0010052 | 14560 | 20 | 278.9252 | 457.2923 | 8.2e-05
0.001 V, 30 Hz |  | 0.001018 | 5220 | 0.00099478 | 0.0010052 | 18050 | 20 | 345.7829 | 1221.9988 | 0.0001578
0.001 V, 40 Hz |  | 0.0010279 | 5110 | 0.00099489 | 0.0010051 | 27850 | 20 | 545.0056 | 271.0393 | 8.61e-05
0.001 V, 55 Hz |  | 0.0010172 | 5110 | 0.00099489 | 0.0010051 | 17150 | 20 | 335.6139 | 1411.5191 | 0.0001864
0.001 V, 300 Hz |  | 0.0010181 | 5110 | 0.00099489 | 0.0010051 | 18080 | 20 | 353.8133 | 418.3983 | 0.000503
0.001 V, 1 KHz |  | 0.0010188 | 5110 | 0.00099489 | 0.0010051 | 18840 | 20 | 368.686 | 485.3259 | 0.000489
0.001 V, 10 KHz |  | 0.0010161 | 5110 | 0.00099489 | 0.0010051 | 16060 | 20 | 314.2833 | 643.5964 | 0.000439
0.001 V, 20 KHz |  | 0.0010172 | 5110 | 0.00099489 | 0.0010051 | 17220 | 20 | 336.9837 | 346.6983 | 0.000556
0.001 V, 30 KHz |  | 0.001019 | 5370 | 0.00099463 | 0.0010054 | 19040 | 20 | 354.5599 | 557.345 | 0.000543
0.001 V, 50 KHz |  | 0.0010204 | 5370 | 0.00099463 | 0.0010054 | 20440 | 20 | 380.6305 | 312.5165 | 0.000414
0.001 V, 100 KHz |  | 0.0010215 | 8900 | 0.0009911 | 0.0010089 | 21530 | 50 | 241.9063 | 373.4515 | 0.000565
0.001 V, 200 KHz |  | 0.0010296 | 16200 | 0.0009838 | 0.0010162 | 29600 | 100 | 182.7126 | 310.8382 | 0.000458
0.001 V, 300 KHz |  | 0.0010163 | 16200 | 0.0009838 | 0.0010162 | 16320 | 100 | 100.7388 | 313.2344 | 0.000355
0.001 V, 500 KHz |  | 0.0010143 | 31700 | 0.0009683 | 0.0010317 | 14270 | 200 | 45.0149 | 560.1397 | 0.000796
0.001 V, 1 MHz |  | 0.0010146 | 33300 | 0.0009667 | 0.0010333 | 14650 | 300 | 43.9922 | 355.4176 | 0.000345
0.01 V, 10 Hz |  | 0.0099998 | 1150 | 0.0099885 | 0.010012 | -18 | 20 | -1.565 | 4.2427 | 3.82e-06
0.01 V, 20 Hz |  | 0.010002 | 820 | 0.0099918 | 0.010008 | 226 | 20 | 27.5528 | 70.4453 | 3.99e-06
0.01 V, 30 Hz |  | 0.010002 | 820 | 0.0099918 | 0.010008 | 203 | 20 | 24.7487 | 51.6199 | 2.31e-05
0.01 V, 40 Hz |  | 0.010002 | 710 | 0.0099929 | 0.010007 | 199 | 20 | 28.0171 | 4.5817 | 2.73e-06
0.01 V, 55 Hz |  | 0.010002 | 710 | 0.0099929 | 0.010007 | 230 | 20 | 32.3815 | 71.7054 | 9.16e-06
0.01 V, 300 Hz |  | 0.010001 | 710 | 0.0099929 | 0.010007 | 110 | 20 | 15.4868 | 15.7921 | 1.197e-05
0.01 V, 1 KHz |  | 0.010001 | 710 | 0.0099929 | 0.010007 | 53 | 20 | 7.4618 | 23.2695 | 6.59e-06
0.01 V, 10 KHz |  | 0.01 | 710 | 0.0099929 | 0.010007 | 18 | 20 | 2.5342 | 14.1507 | 4.47e-06
0.01 V, 20 KHz |  | 0.01 | 710 | 0.0099929 | 0.010007 | 10 | 20 | 1.4079 | 6.4892 | 5.77e-06
0.01 V, 30 KHz |  | 0.0099991 | 970 | 0.0099903 | 0.01001 | -91 | 20 | -9.3794 | 13.5475 | 2e-05
0.01 V, 50 KHz |  | 0.010001 | 970 | 0.0099903 | 0.01001 | 115 | 20 | 11.8532 | 10.0542 | 3.75e-06
0.01 V, 100 KHz |  | 0.010001 | 1700 | 0.009983 | 0.010017 | 60 | 50 | 3.5279 | 6.1234 | 5.51e-06
0.01 V, 200 KHz |  | 0.0099951 | 2700 | 0.009973 | 0.010027 | -494 | 100 | -18.2838 | 25.0246 | 7.13e-06
0.01 V, 300 KHz |  | 0.010001 | 2700 | 0.009973 | 0.010027 | 123 | 100 | 4.5524 | 25.368 | 1.131e-05
0.01 V, 500 KHz |  | 0.010001 | 4700 | 0.009953 | 0.010047 | 135 | 200 | 2.8697 | 57.4857 | 9.89e-06
0.01 V, 1 MHz |  | 0.010001 | 6300 | 0.0099369 | 0.010063 | 58 | 300 | 0.91959 | 10.0741 | 8.91e-06
0.1 V, 10 Hz |  | 0.1 | 710 | 0.099929 | 0.10007 | -5 | 20 | -0.70395 | 2.848 | 5.81e-06
0.1 V, 20 Hz |  | 0.1 | 320 | 0.099968 | 0.10003 | -3 | 20 | -0.93567 | 2.3863 | 2.67e-06
0.1 V, 30 Hz |  | 0.1 | 320 | 0.099968 | 0.10003 | -4 | 20 | -1.2476 | 1.424 | 1.95e-06
0.1 V, 40 Hz |  | 0.1 | 200 | 0.09998 | 0.10002 | -4 | 20 | -1.9901 | 1.0541 | 1.35e-06
0.1 V, 55 Hz |  | 0.1 | 200 | 0.09998 | 0.10002 | 1 | 20 | 0.49752 | 1.9365 | 1.82e-06
0.1 V, 300 Hz |  | 0.1 | 200 | 0.09998 | 0.10002 | -5 | 20 | -2.4876 | 1.1667 | 1.71e-06
0.1 V, 1 KHz |  | 0.1 | 200 | 0.09998 | 0.10002 | -3 | 20 | -1.4926 | 1.1667 | 1.51e-06
0.1 V, 10 KHz |  | 0.1 | 200 | 0.09998 | 0.10002 | -2 | 20 | -0.99504 | 0.5 | 7e-07
0.1 V, 20 KHz |  | 0.1 | 200 | 0.09998 | 0.10002 | 3 | 20 | 1.4926 | 0.7071 | 5.6e-07
0.1 V, 30 KHz |  | 0.1 | 430 | 0.099957 | 0.10004 | 4 | 20 | 0.92923 | 0.83333 | 8.6e-07
0.1 V, 50 KHz |  | 0.1 | 430 | 0.099957 | 0.10004 | 4 | 20 | 0.92923 | 1.1667 | 4.3e-07
0.1 V, 100 KHz |  | 0.1 | 1100 | 0.09989 | 0.10011 | 11 | 50 | 0.99897 | 0.66666 | 5.8e-07
0.1 V, 200 KHz |  | 0.10004 | 1300 | 0.09987 | 0.10013 | 417 | 100 | 31.9824 | 1.1299 | 3.4e-07
0.1 V, 300 KHz |  | 0.1 | 1300 | 0.09987 | 0.10013 | 30 | 100 | 2.3009 | 1.2018 | 4.7e-07
0.1 V, 500 KHz |  | 0.10001 | 2100 | 0.099789 | 0.10021 | 65 | 200 | 3.0813 | 1.3016 | 1.01e-06
0.1 V, 1 MHz |  | 0.1 | 4300 | 0.099569 | 0.10043 | 38 | 300 | 0.88158 | 4.7724 | 8.3e-07
1 V, 10 Hz |  | 0.99999 | 650 | 0.99935 | 1.0007 | -9 | 10 | -1.3845 | 2.1279 | 5.42e-06
1 V, 20 Hz |  | 1 | 200 | 0.9998 | 1.0002 | 0 | 10 | 0 | 2.4381 | 2.12e-06
1 V, 30 Hz |  | 1 | 200 | 0.9998 | 1.0002 | -4 | 10 | -1.9975 | 0.72649 | 2e-06
1 V, 40 Hz |  | 1 | 82 | 0.99992 | 1.0001 | -4 | 10 | -4.8422 | 0.88192 | 1.55e-06
1 V, 55 Hz |  | 1 | 82 | 0.99992 | 1.0001 | -1 | 10 | -1.2105 | 1.1667 | 2.1e-06
1 V, 300 Hz |  | 1 | 82 | 0.99992 | 1.0001 | 0 | 10 | 0 | 2 | 1.91e-06
1 V, 1 KHz |  | 1 | 82 | 0.99992 | 1.0001 | -2 | 10 | -2.4211 | 1.3229 | 9.9e-07
1 V, 10 KHz |  | 0.99999 | 82 | 0.99992 | 1.0001 | -5 | 10 | -6.0527 | 0.5 | 5.1e-07
1 V, 20 KHz |  | 0.99999 | 82 | 0.99992 | 1.0001 | -5 | 10 | -6.0527 | 0.66667 | 6.9e-07
1 V, 30 KHz |  | 0.99999 | 150 | 0.99985 | 1.0002 | -5 | 10 | -3.326 | 0.86603 | 3e-07
1 V, 50 KHz |  | 1 | 150 | 0.99985 | 1.0002 | -3 | 20 | -1.9825 | 0.66667 | 5.7e-07
1 V, 100 KHz |  | 1 | 330 | 0.99967 | 1.0003 | -2 | 30 | -0.60357 | 0.97183 | 1.11e-06
1 V, 200 KHz |  | 1.0002 | 590 | 0.99941 | 1.0006 | 175 | 70 | 29.4544 | 1.1178 | 4.3e-07
1 V, 300 KHz |  | 1 | 590 | 0.99941 | 1.0006 | 6 | 70 | 1.0099 | 0.5 | 4.8e-07
1 V, 500 KHz |  | 1 | 1500 | 0.9985 | 1.0015 | 13 | 100 | 0.86475 | 0.7071 | 3.1e-07
1 V, 1 MHz |  | 1.0001 | 3200 | 0.99679 | 1.0032 | 72 | 200 | 2.2456 | 1.4813 | 4.5e-07
10 V, 10 Hz |  | 10 | 650 | 9.9935 | 10.0065 | -3 | 10 | -0.46148 | 2.4037 | 4.01e-06
10 V, 20 Hz |  | 9.9999 | 200 | 9.998 | 10.002 | -6 | 10 | -2.9963 | 2.5981 | 2.25e-06
10 V, 30 Hz |  | 9.9999 | 200 | 9.998 | 10.002 | -9 | 10 | -4.4944 | 1.5366 | 1.84e-06
10 V, 40 Hz |  | 10 | 82 | 9.9992 | 10.0008 | -5 | 10 | -6.0527 | 1.0138 | 1.61e-06
10 V, 55 Hz |  | 10 | 82 | 9.9992 | 10.0008 | -2 | 10 | -2.4211 | 1.9003 | 2.25e-06
10 V, 300 Hz |  | 10 | 82 | 9.9992 | 10.0008 | -4 | 10 | -4.8422 | 2.7588 | 3.2e-06
10 V, 1 KHz |  | 10 | 82 | 9.9992 | 10.0008 | -5 | 10 | -6.0527 | 1.1304 | 8.1e-07
10 V, 10 KHz |  | 9.9999 | 82 | 9.9992 | 10.0008 | -6 | 10 | -7.2633 | 0.78174 | 2.6e-07
10 V, 20 KHz |  | 10 | 82 | 9.9992 | 10.0008 | -4 | 10 | -4.8422 | 0.78174 | 6.1e-07
10 V, 30 KHz |  | 10 | 150 | 9.9985 | 10.0015 | -5 | 10 | -3.326 | 0.66667 | 8e-07
10 V, 50 KHz |  | 10 | 150 | 9.9985 | 10.0015 | -4 | 20 | -2.6433 | 0.70711 | 6.3e-07
10 V, 100 KHz |  | 10 | 290 | 9.9971 | 10.0029 | -4 | 30 | -1.372 | 0.70711 | 6e-07
10 V, 200 KHz |  | 10.0008 | 720 | 9.9928 | 10.0072 | 79 | 70 | 10.9207 | 0.60088 | 4.3e-07
10 V, 300 KHz |  | 10 | 720 | 9.9928 | 10.0072 | -5 | 70 | -0.69119 | 0.86603 | 5.4e-07
10 V, 500 KHz |  | 10 | 1800 | 9.982 | 10.018 | 5 | 100 | 0.27735 | 0.66666 | 4.1e-07
10 V, 1 MHz |  | 10.0002 | 3700 | 9.9629 | 10.0371 | 24 | 200 | 0.6477 | 1.1666 | 8.6e-07
19 V, 10 Hz |  | 19.0022 | 602.63 | 18.9885 | 19.0115 | 113.6842 | 10 | 18.8621 | 2.2804 | 4.28e-06
19 V, 20 Hz |  | 19 | 185.79 | 18.9965 | 19.0035 | 0 | 10 | 0 | 1.2621 | 2.09e-06
19 V, 30 Hz |  | 18.9995 | 185.79 | 18.9965 | 19.0035 | -25.2632 | 10 | -13.578 | 0.61405 | 1.93e-06
19 V, 40 Hz |  | 18.9994 | 78.684 | 18.9985 | 19.0015 | -28.9474 | 10 | -36.4958 | 0.69627 | 1.98e-06
19 V, 55 Hz |  | 18.9994 | 78.684 | 18.9985 | 19.0015 | -33.1579 | 10 | -41.8043 | 1.9713 | 2.13e-06
19 V, 300 Hz |  | 18.9992 | 78.684 | 18.9985 | 19.0015 | -44.7368 | 10 | -56.4027 | 1.256 | 2.03e-06
19 V, 1 KHz |  | 18.9999 | 78.684 | 18.9985 | 19.0015 | -5.7895 | 10 | -7.2992 | 1.2467 | 1.32e-06
19 V, 10 KHz |  | 18.9991 | 78.684 | 18.9985 | 19.0015 | -46.8421 | 10 | -59.0569 | 1.0342 | 4e-07
19 V, 20 KHz |  | 18.9992 | 78.684 | 18.9985 | 19.0015 | -44.7368 | 10 | -56.4027 | 0.52634 | 4.3e-07
19 V, 30 KHz |  | 18.9991 | 140.53 | 18.9973 | 19.0027 | -46.3158 | 10 | -32.8748 | 0.26317 | 1e-07
19 V, 50 KHz |  | 18.9992 | 140.53 | 18.9973 | 19.0027 | -43.1579 | 20 | -30.4044 | 0.2774 | 1.9e-07
19 V, 100 KHz |  | 18.9994 | 271.05 | 18.9948 | 19.0052 | -32.1053 | 30 | -11.7729 | 0.38237 | 2.9e-07
19 V, 200 KHz |  | 19.0016 | 639.47 | 18.9878 | 19.0122 | 83.6842 | 70 | 13.0088 | 0.82281 | 4e-07
19 V, 300 KHz |  | 19.0005 | 639.47 | 18.9878 | 19.0122 | 25.2632 | 70 | 3.9272 | 0 | 3.8e-07
19 V, 500 KHz |  | 19.0019 | 1563.2 | 18.9702 | 19.0298 | 101.0526 | 100 | 6.4513 | 0.68504 | 4e-07
19 V, 1 MHz |  | 19.0088 | 3273.7 | 18.9377 | 19.0623 | 464.2105 | 200 | 14.1536 | 12.0264 | 4.3e-07
100 V, 10 Hz |  | 100 | 650 | 99.935 | 100.065 | 0 | 10 | 0 | 4.0654 | 4.81e-06
100 V, 20 Hz |  | 99.9996 | 200 | 99.98 | 100.02 | -4 | 10 | -1.9975 | 1.0138 | 1.91e-06
100 V, 30 Hz |  | 99.9996 | 200 | 99.98 | 100.02 | -4 | 10 | -1.9975 | 0.88192 | 2.16e-06
100 V, 40 Hz |  | 99.9996 | 90 | 99.9909 | 100.0091 | -4 | 10 | -4.4173 | 1.4142 | 1.64e-06
100 V, 55 Hz |  | 99.9995 | 90 | 99.9909 | 100.0091 | -5 | 10 | -5.5216 | 1.1304 | 2.01e-06
100 V, 300 Hz |  | 99.9994 | 90 | 99.9909 | 100.0091 | -6 | 10 | -6.6259 | 2.2048 | 2.62e-06
100 V, 1 KHz |  | 99.9996 | 90 | 99.9909 | 100.0091 | -4 | 10 | -4.4173 | 0.72649 | 1.68e-06
100 V, 10 KHz |  | 99.9995 | 90 | 99.9909 | 100.0091 | -5 | 10 | -5.5216 | 0.5 | 5e-07
100 V, 20 KHz |  | 99.9995 | 90 | 99.9909 | 100.0091 | -5 | 10 | -5.5216 | 0.60093 | 4.9e-07
100 V, 30 KHz |  | 99.9995 | 260 | 99.974 | 100.026 | -5 | 10 | -1.9217 | 0.66667 | 7.5e-07
100 V, 50 KHz |  | 99.9994 | 260 | 99.9739 | 100.0261 | -6 | 20 | -2.3009 | 1.236 | 6e-07
100 V, 100 KHz |  | 99.9995 | 650 | 99.9349 | 100.0651 | -5 | 30 | -0.76841 | 1.6915 | 5.8e-07
100 V, 200 KHz |  | 99.9995 | 2600 | 99.74 | 100.26 | -5 | 50 | -0.19227 | 3.8079 | 1.7e-06
1000 V, 55 Hz |  | 999.993 | 84 | 999.9147 | 1000.0853 | -7 | 15 | -8.2036 | 1.5811 | 2.06e-06
1000 V, 300 Hz |  | 999.99 | 84 | 999.9147 | 1000.0853 | -10 | 15 | -11.7194 | 2.2236 | 2.4e-06
700 V, 1 KHz |  | 700.005 | 85.714 | 699.9305 | 700.0695 | 7.1429 | 50 | 7.1982 | 1.5058 | 1.62e-06
DCI PERFORMANCE TEST
DCI Test | Expected Value | Measured A | Source Unc. ppm | Lower Limit A | Upper Limit A | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
1 µADC | 1e-06 | 1.0006e-06 | 10043.4717 |  |  | 650 |  |  | 57.6975 | 1.668e-05
10 µADC | 1e-05 | 1.0001e-05 | 1049.937 |  |  | 60 |  |  | 4.9997 | 2.65e-06
50 µADC | 5e-05 | 5e-05 | 250 |  |  | 5 |  |  | 1.9148 | 4.1e-07
100 µADC | 0.0001 | 0.0001 | 150 | 9.9985e-05 | 0.00010002 | -0.5 | 7 | -0.33297 | 0.57735 | 2.4e-07
-100 µADC | -0.0001 | -9.9999e-05 | 150 | -9.9985e-05 | -0.00010002 | -5 | 7 | -3.3297 | -0.5 | 2.5e-07
-50 µADC | -5e-05 | -4.9999e-05 | 250 |  |  | -14 |  |  | -2 | 4.9e-07
-10 µADC | -1e-05 | -9.9993e-06 | -1050.0735 |  |  | -70 |  |  | -5.0004 | 2.09e-06
0.1 mADC | 0.0001 | 9.9998e-05 | 150 |  |  | -20 |  |  | 0 | 1.65e-06
0.5 mADC | 0.0005 | 0.0005 | 70 |  |  | 4 |  |  | 0 | 3.3e-07
1 mADC | 0.001 | 0.001 | 60 | 0.00099994 | 0.0010001 | 1 | 7 | 1.6554 | 0.5 | 1.7e-07
-1 mADC | -0.001 | -0.00099999 | 60 | -0.00099994 | -0.0010001 | -7 | 7 | -11.5881 | -0.5 | 1.5e-07
-0.5 mADC | -0.0005 | -0.00049999 | 70 |  |  | -12 |  |  | -1 | 2.9e-07
1 mADC | 0.001 | 0.001 | 60 |  |  | 5 |  |  | 5.7735 | 1.72e-06
5 mADC | 0.005 | 0.005 | 70 |  |  | 6 |  |  | 0.99999 | 4.3e-07
10 mADC | 0.01 | 0.01 | 60 | 0.0099994 | 0.010001 | 3 | 7 | 4.9663 | 0 | 2.1e-07
-10 mADC | -0.01 | -0.01 | 60 | -0.0099994 | -0.010001 | -4 | 7 | -6.6218 | -0.5 | 1.1e-07
-5 mADC | -0.005 | -0.005 | 70 |  |  | -8 |  |  | -1 | 2.9e-07
10 mADC | 0.01 | 0.01 | 60 |  |  | 30 |  |  | 4.9999 | 1.56e-06
50 mADC | 0.05 | 0.050001 | 80 |  |  | 20 |  | 25 | 0.99998 | 2.6e-07
100 mADC | 0.1 | 0.1 | 70 | 0.099993 | 0.10001 | 15 | 7 | 21.3222 | 0.49999 | 1.8e-07
-100 mADC | -0.1 | -0.1 | 70 | -0.099993 | -0.10001 | 6.5 | 7 | 9.2396 | -0.57735 | 1.3e-07
-50 mADC | -0.05 | -0.05 | 80 |  |  | 2 |  |  | -1 | 2.9e-07
0.1 ADC | 0.1 | 0.1 | 70 |  |  | 35 |  |  | 5.7733 | 1.96e-06
0.5 ADC | 0.5 | 0.50002 | 140 |  |  | 40 |  |  | 0.99996 | 2.9e-07
-0.5 ADC | -0.5 | -0.5 | 140 |  |  | 4 |  |  | -1 | 3.4e-07
-1.0 ADC | -1 | -0.99997 | 110 | -0.99989 | -1.0001 | -28 | 20 | -25.044 | 0 | 1.9e-07
1.0 ADC | 1 | 0.99999 | 110 | 0.99989 | 1.0001 | -14.5 | 20 | -12.9692 | 0.57736 | 1.9e-07
ACI PERFORMANCE TEST
ACI Test | Expected Value | Measured Aac | Source Unc. ppm | Lower Limit Aac | Upper Limit Aac | Measured ppm | DUT Transfer STB ppm | Test Result % of SPEC | MEAS SDEV ppm
10 µA, 10 Hz | 1e-05 | 1.0015e-05 | 3694.5874 |  |  | 1465 |  |  | 5.7651 | 6.51e-06
10 µA, 20 Hz | 1e-05 | 1.0015e-05 | 2875.6578 |  |  | 1510 |  |  | 4.9925 | 3.03e-06
10 µA, 30 Hz | 1e-05 | 1.0015e-05 | 2875.6434 |  |  | 1515 |  |  | 9.5598 | 4.42e-06
10 µA, 40 Hz | 1e-05 | 1.0015e-05 | 2136.7628 |  |  | 1515 |  |  | 9.5598 | 3.15e-06
10 µA 55 Hz | 1e-05 | 1.0015e-05 | 2136.6988 |  |  | 1545 |  |  | 9.5595 | 5.59e-06
10 µA, 300 Hz | 1e-05 | 1.0016e-05 | 2136.5708 |  |  | -40 |  |  | 9.5589 | 5.58e-06
10 µA, 1 KHz | 1e-05 | 1.0016e-05 | 2136.5388 |  |  | 1620 |  |  | 4.9919 | 9.45e-06
10 µA, 5 KHz | 1e-05 | 1.0015e-05 | 5591.6405 |  |  | 1495 |  |  | 14.121 | 1.466e-05
10 µA, 10 KHz | 1e-05 | 1.0011e-05 | 11587.1961 |  |  | 1105 |  |  | 9.5637 | 2.13e-05
100 µA, 10 Hz | 0.0001 | 9.9998e-05 | 1000 | 9.9939e-05 | 0.00010006 | -40 | 50 | -3.995 | 4.933 | 3.12e-06
100 µA, 20 Hz | 0.0001 | 9.9999e-05 | 630 | 9.9937e-05 | 0.00010006 | -5.5 | 50 | -0.87028 | 2.5 | 1e-06
100 µA, 30 Hz | 0.0001 | 0.0001 | 630 | 9.9939e-05 | 0.00010006 | -4 | 50 | -0.63293 | 6.9462 | 6.1e-07
100 µA, 40 Hz | 0.0001 | 9.9999e-05 | 340 | 9.9966e-05 | 0.00010003 | -7.5 | 50 | -2.1824 | 4.0825 | 2.39e-06
100 µA 55 Hz | 0.0001 | 9.9999e-05 | 340 | 9.9939e-05 | 0.00010006 | -40 | 50 | -11.6395 | 2.3629 | 1.7e-06
100 µA, 300 Hz | 0.0001 | 9.9999e-05 | 340 | 9.9966e-05 | 0.00010003 | -7 | 50 | -2.0369 | 0.5 | 2.67e-06
100 µA, 1 KHz | 0.0001 | 9.9999e-05 | 340 | 9.9939e-05 | 0.00010006 | -7 | 50 | -2.0369 | 3.2016 | 2.32e-06
100 µA, 5 KHz | 0.0001 | 9.9998e-05 | 0.11 | 9.999e-05 | 0.00010001 | -20 | 100 | -20 | 1.1547 | 1.1e-06
100 µA, 10 KHz | 0.0001 | 9.9996e-05 | 0.26 | 9.9939e-05 | 0.00010006 | -35 | 300 | -11.6667 | 1.5001 | 9.5e-07
1mA, 10 Hz | 0.001 | 0.001 | 750 | 0.00099925 | 0.0010008 | -2 | 40 | -0.26629 | 3.304 | 4.5e-06
1 mA, 20 Hz | 0.001 | 0.00099999 | 420 | 9.9939e-05 | 0.00010006 | -5.5 | 40 | -1.3036 | 1.291 | 2.04e-06
1 mA, 30 Hz | 0.001 | 0.001 | 420 | 0.00099958 | 0.0010004 | -2.5 | 40 | -0.59256 | 6.7823 | 1.8e-07
1 mA, 40 Hz | 0.001 | 0.001 | 180 | 9.9939e-05 | 0.00010006 | -4.5 | 40 | -2.4405 | 0.57735 | 2.48e-06
1 mA, 55 Hz | 0.001 | 0.001 | 180 | 0.00099982 | 0.0010002 | -3.5 | 40 | -1.8981 | 3.5 | 1.95e-06
1 mA, 300 Hz | 0.001 | 0.001 | 180 | 9.9939e-05 | 0.00010006 | -0.5 | 40 | -0.27116 | 4.0415 | 1.93e-06
1 mA, 1 KHz | 0.001 | 0.00099999 | 180 | 0.00099982 | 0.0010002 | -6 | 40 | -3.254 | 0.8165 | 1.2e-06
1 mA, 5 KHz | 0.001 | 0.00099999 | 0.11 | 9.9939e-05 | 0.00010006 | -12 | 70 | -17.1428 | 0.50001 | 4.2e-07
1 mA, 10 KHz | 0.001 | 0.00099997 | 0.26 | 0.0009998 | 0.0010002 | -25 | 200 | -12.5 | 0.81652 | 7.3e-07
10 mA, 10 Hz | 0.01 | 0.01 | 750 | 9.9939e-05 | 0.00010006 | 3.5 | 40 | 0.466 | 5.3229 | 5.65e-06
10 mA, 20 Hz | 0.01 | 0.01 | 420 | 0.0099958 | 0.010004 | 1 | 40 | 0.23702 | 2.2174 | 1.01e-06
10 mA, 30 Hz | 0.01 | 0.01 | 420 | 9.9939e-05 | 0.00010006 | -1 | 40 | -0.23702 | 7.8475 | 1.05e-06
10 mA, 40 Hz | 0.01 | 0.01 | 180 | 0.0099982 | 0.010002 | -2.5 | 40 | -1.3558 | 1.4142 | 2.44e-06
10 mA, 55 Hz | 0.01 | 0.01 | 180 | 9.9939e-05 | 0.00010006 | 0 | 40 | 0 | 0.8165 | 2.12e-06
10 mA, 300 Hz | 0.01 | 0.0099999 | 180 | 0.0099982 | 0.010002 | -8 | 40 | -4.3386 | 1.2583 | 3.25e-06
10 mA, 1 KHz | 0.01 | 0.0099999 | 180 | 9.9939e-05 | 0.00010006 | -10.5 | 40 | -5.6944 | 0.57736 | 8.8e-07
10 mA, 5 KHz | 0.01 | 0.0099996 | 0.11 | 0.0099993 | 0.010001 | -37.5 | 70 | -53.5714 | 1.291 | 6e-07
10 mA, 10 KHz | 0.01 | 0.0099996 | 0.26 | 9.9939e-05 | 0.00010006 | -42 | 200 | -21 | 0.50002 | 3.6e-07
100 mA, 10 Hz | 0.1 | 0.1 | 750 | 0.099925 | 0.10008 | -1 | 40 | -0.13314 | 2.9861 | 4.1e-06
100 mA, 20 Hz | 0.1 | 0.1 | 420 | 9.9939e-05 | 0.00010006 | -2 | 40 | -0.47405 | 2.4495 | 9.6e-07
100 mA, 30 Hz | 0.1 | 0.1 | 420 | 0.099958 | 0.10004 | -1.5 | 40 | -0.35553 | 6.0759 | 7.7e-07
100 mA, 40 Hz | 0.1 | 0.1 | 190 | 9.9939e-05 | 0.00010006 | -0.5 | 40 | -0.25751 | 0.57735 | 2.52e-06
100 mA, 55 Hz | 0.1 | 0.1 | 190 | 0.099981 | 0.10002 | 0 | 40 | 0 | 1.5 | 1.36e-06
100 mA, 300 Hz | 0.1 | 0.1 | 190 | 9.9939e-05 | 0.00010006 | -2 | 40 | -1.0301 | 1.633 | 2.4e-06
100 mA, 1 KHz | 0.1 | 0.099999 | 190 | 0.099981 | 0.10002 | -6.5 | 40 | -3.3477 | 0.57735 | 1.24e-06
100 mA, 5 KHz | 0.1 | 0.099999 | 0.11 | 9.9939e-05 | 0.00010006 | -10.5 | 70 | -15 | 0.57736 | 5.9e-07
100 mA, 10 KHz | 0.1 | 0.099998 | 0.26 | 0.09998 | 0.10002 | -16.5 | 200 | -8.25 | 0.57736 | 4.6e-07
1 A, 10 Hz | 1 | 1 | -1 | 9.9939e-05 | 0.00010006 | 19 | 40 | 47.4852 | 2.9439 | 3.36e-06
1 A, 20 Hz | 1 | 1 | -1 | 0.99996 | 1 | 16 | 40 | 39.9875 | 3.0956 | 1.18e-06
1 A, 30 Hz | 1 | 1 | -1 | 9.9939e-05 | 0.00010006 | 5.5 | 40 | 13.7457 | 5.1234 | 2.5e-06
1 A, 40 Hz | 1 | 1 | 690 | 0.99931 | 1.0007 | 9 | 40 | 1.3022 | 1.5 | 2.43e-06
1 A, 55 Hz | 1 | 1 | 690 | 9.9939e-05 | 0.00010006 | 11.5 | 40 | 1.6639 | 1.291 | 2.37e-06
1 A, 300 Hz | 1 | 1 | 690 | 0.99931 | 1.0007 | -0.5 | 40 | -0.072342 | 1.7078 | 2.02e-06
1 A, 1 KHz | 1 | 1 | 690 | 9.9939e-05 | 0.00010006 | 6 | 40 | 0.86811 | 0.5 | 1.59e-06
1 A, 5 KHz | 1 | 1 | 850 | 0.99915 | 1.0009 | 6 | 70 | 0.7035 | 0.5 | 5.8e-07
1 A, 10 KHz | 1 | 1 | 0.92 | 9.9939e-05 | 0.00010006 | 24 | 200 | 11.9999 | 0.49999 | 3.1e-07
Test Summary | Initialization | DCV | OHM | ACV | DCI | ACI |  | Total
Time escape minute | 0.19 | 206.94 | 262.31 | 238.26 | 40.42 | 176.02 |  | 924.14